# ANSYS HFSS Student Edition Model Construction Guide

## Overview

This document provides the complete procedure for constructing, simulating, and post-processing the proposed HFSS model using Ansys Electronics Desktop (HFSS Student Edition). The model is designed for a 60 mm × 60 mm tissue footprint and is optimized to remain within the mesh limitations of the Student Edition while allowing a parametric sweep of subcutaneous fat thickness.

# Phase 1 – Environment Setup & Variable Mapping

## 1.1: Create a New Project

- Open Ansys Electronics Desktop.

- Create a New HFSS Design.

## 1.2: Define Design Variables

Navigate to:

HFSS → Design Properties → Add

Create the following local variables.

| Variable | Value | Description |
| --- | --- | --- |
| fat_thickness | 5 mm | Baseline variable for parametric sweep |
| muscle_thickness | 15 mm | Muscle layer thickness |
| skin_thickness | 1.5 mm | Skin layer thickness |
| tissue_width | 60 mm | Overall tissue width and length |
| superstrate_thickness | 1 mm | Superstrate thickness |
| patch_length | 14 mm | Patch antenna length |
| patch_width | 14 mm | Patch antenna width |

## 1.3: Set Solution Type

Right-click the design name in the Project Tree.

Select:

Solution Type → Terminal

# Phase 2 – Biological Tissue Model Construction

Construct three stacked biological tissue layers.

## 2.1 Muscle Layer

Create a rectangular box.

### Geometry

| Parameter | Value |
| --- | --- |
| Material | Muscle (or Custom) |
| Relative Permittivity (εr) | 54.5 |
| Conductivity (σ) | 0.94 S/m |
| Position | (-tissue_width/2, -tissue_width/2, 0 mm) |
| X Size | tissue_width |
| Y Size | tissue_width |
| Z Size | muscle_thickness |

Evaluated position:

- X = −30 mm

- Y = −30 mm

- Z = 0 mm

## 2.2 Subcutaneous Fat Layer

Create a second box directly above the muscle.

### Geometry

| Parameter | Value |
| --- | --- |
| Material | Fat (or Custom) |
| Relative Permittivity (εr) | 5.3 |
| Conductivity (σ) | 0.04 S/m |
| Position | (-tissue_width/2, -tissue_width/2, muscle_thickness) |
| X Size | tissue_width |
| Y Size | tissue_width |
| Z Size | fat_thickness |

Evaluated position:

- X = −30 mm

- Y = −30 mm

- Z = 15 mm

This layer will be varied parametrically.

## 2.3 Skin Layer

Create the third biological layer.

### Geometry

| Parameter | Value |
| --- | --- |
| Material | Skin (or Custom) |
| Relative Permittivity (εr) | 41.4 |
| Conductivity (σ) | 0.69 S/m |
| Position | (-tissue_width/2, -tissue_width/2, muscle_thickness + fat_thickness) |
| X Size | tissue_width |
| Y Size | tissue_width |
| Z Size | skin_thickness |

Baseline evaluated position:

- X = −30 mm

- Y = −30 mm

- Z = 20 mm

# Phase 3 – Structural Component Modeling

## 3.1 Ground Plane

Create a rectangle.

### Geometry

| Parameter | Value |
| --- | --- |
| Position | (-tissue_width/2, -tissue_width/2, 0 mm) |
| Axis | Z |
| X Size | tissue_width |
| Y Size | tissue_width |

Assign Boundary:

Perfect E

## 3.2 Internal Sensor Cylinder

Create a cylinder.

### Geometry

| Parameter | Value |
| --- | --- |
| Center Position | (0, 0, 17.5 mm) |
| Axis | Z |
| Radius | 0.5 mm |
| Height | 2 mm |
| Number of Segments | 0 |

This places the sensor approximately at the center of the fat layer.

## 3.3 External Patch Antenna

Create a rectangle.

### Geometry

| Parameter | Value |
| --- | --- |
| Position | (-patch_length/2, -patch_width/2, muscle_thickness + fat_thickness + skin_thickness) |
| Axis | Z |
| X Size | patch_length |
| Y Size | patch_width |

Baseline evaluated position:

- X = −7 mm

- Y = −7 mm

- Z = 21.5 mm

Assign Boundary:

Perfect E

## 3.4 Ceramic Superstrate

Create a dielectric box.

### Geometry

| Parameter | Value |
| --- | --- |
| Material | Alumina (or Custom Ceramic) |
| Position | (-30 mm, -30 mm, muscle_thickness + fat_thickness + skin_thickness) |
| X Size | 60 mm |
| Y Size | 60 mm |
| Z Size | 2 mm |

Baseline evaluated position:

- X = −30 mm

- Y = −30 mm

- Z = 21.5 mm

## 3.5 Lumped Port Feed Bridge

Create a rectangle.

### Geometry

| Parameter | Value |
| --- | --- |
| Position | (0, 0, muscle_thickness + fat_thickness/2) |
| Axis | Y |
| Width (X Size) | 0.5 mm |
| Height (Z Size) | -(muscle_thickness + fat_thickness/2) |

Baseline evaluated values:

- Feed Position = (0, 0, 17.5 mm)

- Feed Extension = −17.5 mm

Assign Excitation:

Lumped Port

Reference Impedance:

50 Ω

## 3.6 Radiation Region

Create the simulation region.

### Region Settings

Padding Type:

Absolute Offset

Padding in all six directions:

248.13895781638 mm

Assign Boundary:

Radiation

This corresponds approximately to a quarter-wavelength clearance at 600 MHz.

# Phase 4 – Analysis and Parametric Sweep

## 4.1 Solution Setup

Navigate to:

Analysis → Add Solution Setup

Use the following settings.

| Parameter | Value |
| --- | --- |
| Center Frequency | 600 MHz |
| Maximum Passes | 10 |
| Maximum Delta S | 0.02 |

These settings are optimized for the Student Edition mesh limitations.

## 4.2 Frequency Sweep

Right-click the created setup.

Select:

Add Frequency Sweep

Use the following configuration.

| Parameter | Value |
| --- | --- |
| Sweep Type | Fast (or Interpolating) |
| Start Frequency | 200 MHz |
| Stop Frequency | 800 MHz |
| Step Size | 1 MHz |

This produces 601 frequency samples.

## 4.3 Parametric Sweep

Navigate to:

Optimetrics → Add → Parametric

Choose the variable:

fat_thickness

Configure the sweep.

| Parameter | Value |
| --- | --- |
| Sweep Type | Linear Step |
| Start | 5 mm |
| Stop | 15 mm |
| Step | 2.5 mm |

The simulation automatically evaluates five tissue configurations:

- 5 mm

- 7.5 mm

- 10 mm

- 12.5 mm

- 15 mm

Under the Options tab, enable:

Save Fields and Mesh

# Phase 5 – Validation and Simulation

## Validation

From the toolbar select:

HFSS → Validation Check

Verify that:

- No geometry errors exist.

- Material assignments are complete.

- Boundary conditions are valid.

- Port definitions are correct.

- Radiation region is properly assigned.

## Run Simulation

Navigate to:

Optimetrics → Parametric Setup

Right-click and select:

Analyze

HFSS will execute all parameter combinations automatically.

# Phase 6 – Post-Processing and Result Generation

After the simulation completes, generate the following engineering plots.

## 6.1 S-Parameter Plot

Navigate to:

Results → Create Terminal Solution Data Report → Rectangular Plot

Configuration:

| Setting | Value |
| --- | --- |
| Category | Terminal S Parameter |
| Quantity | S(1,1) |
| Function | dB |

Create the report.

## 6.2 VSWR Plot

Navigate to:

Results → Create Terminal Solution Data Report → Rectangular Plot

Configuration:

| Setting | Value |
| --- | --- |
| Category | Terminal VSWR |
| Quantity | VSWR(1) |
| Function | None |

Create the report.

## 6.3 Input Impedance Plot

Navigate to:

Results → Create Terminal Solution Data Report → Rectangular Plot

### Real Component

| Setting | Value |
| --- | --- |
| Category | Terminal Z Parameter |
| Quantity | Z(1,1) |
| Function | Re |

Click Add Trace.

### Imaginary Component

Use the same settings except:

| Setting | Value |
| --- | --- |
| Function | Im |

Create the report containing both traces.

## 6.4 2D Far-Field Gain Plot

Navigate to:

Results → Create Far Fields Report → Rectangular Plot

Configuration:

| Setting | Value |
| --- | --- |
| Category | Gain |
| Quantity | GainTotal |
| Function | dB |

Under the Families tab, select the desired angular cut:

Create the report.

## 6.5 3D Radiation Pattern

Navigate to:

Results → Create Far Fields Report → 3D Polar Plot

Configuration:

| Setting | Value |
| --- | --- |
| Category | Gain |
| Quantity | GainTotal |
| Function | dB |

Create the report to generate the complete three-dimensional radiation pattern.

# Expected Outputs

The simulation should generate the following result set:

- S11 (Return Loss) vs Frequency

- VSWR vs Frequency

- Input Resistance (Real Impedance)

- Input Reactance (Imaginary Impedance)

- 2D Far-Field Gain Pattern

- 3D Radiation Pattern

Additionally, the parametric sweep will evaluate all five fat-layer thickness configurations, enabling comparison of antenna performance as the subcutaneous fat thickness varies.